## S01 Protocol Sheet
Flow Cytometry Protocol Sheet S01 - BD FACSymphony S6 RTP 2104 919-684-4130
Please Schedule Job in CoreResearch, Paste Reservation ID below, and email Protocol Sheet to dhviflo@dm.duke.edu
REQUIRED:
Reservation ID
Registered Lab PI/EndUser: | [7 empty] | # of Events/Cells to Record: |  |  |  |  | Please acknowledge the DHVI Flow Facility in abstracts and publications: “Flow cytometry was performed in the Duke Human Vaccine Institute Research Flow Cytometry Shared Resource Facility (Durham, NC).”
End UserTelephone/Email: | [7 empty] | Record Gate:
Hazard Review: | [7 empty] | Duke Box Data Folder:
Hazard Assessment Number:
Assigned BSL Containment Level:
Sort Pressure (70, 45, 20 psi):
Cell Type: | [7 empty] | Sort Nozzle (70, 85, 100 µm):
Cell Line Name (if applicable): | [7 empty] | Sort Setup (1-4 way, tubes, plates, etc.):
Sample Prep Date: | [7 empty] | Sort Precision Mode (single cell_PH2, single cell, purity etc.):
Custom Setup (index sort, chiller, neutral density filter, etc.):
SPICA Number (if applicable):
Comments:
OPTIONAL: | Laser | Ultraviolet, 355 nm, 65 mW | [7 empty] | Violet, 402 nm, 200 mW | [7 empty] | Blue, 488 nm, 200 mW | [5 empty] | Green, 561 nm, 150 mW |  |  |  |  | Red, 638 nm, 150 mW
| Filter | UV-405LP_379_28 | UV-410LP_450_50 | UV-450LP_515_30 | UV-550LP_586_15 | UV-600LP_610_20 | UV-635LP_670_30 | UV-690LP_740_35 | UV-770LP_820_60 | V-410LP_431_28 | V-450LP_470_14 | V-550LP_586_15 | V-600LP_610_20 | V-635LP_670_30 | V-690LP_710_50 | V-740LP_750_30 | V-750LP_780_60 | B-505LP_515_20 | B-600LP_610_20 | B-635LP_670_30 | B-690LP_710_50 | B-730LP_750_30 | B-770LP_780_60 | G-570LP_586_15 | G-600LP_610_20 | G-650LP_670_30 | G-690LP_710_50 | G-750LP_780_60 | R-655LP_670_30 | R-690LP_710_50 | R-750LP_780_60
Expt PMT Voltage
| Common Fluor(s) | BUV396 | DAPI | BUV496 | BUV563 | BUV615 | BUV661 | BUV737 | BUV805 | BV421 | BV480 | BV570 | BV605 | BV650 | BV711 | BV750 | BV786 | BB515 | BB630 | BB660 | BB700 | BB750 | BB790 | PE | PE-CF594 | PE-Cy5 | PE-Cy5.5 | PE-Cy7 | APC | APC-Cy5.5 | APC-Cy7
REQUIRED: Insert additional rows as needed for more tubes. Number consecutively.
Sample ID | Tube # | Please enter the name of the actual Fluors you are using
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
Facility Operator Use:
| Acquisition Date | Start Time | End Time | Hours | Operator Initials | Add. Operator | Additional Hours | Initials | Total Hours | Scheduled Time
|  |  | TOTAL | 0 |  |  | 0 |  | 0
Sheath Pressure (psi) | Nozzle Size (µM) | Nozzle Number | Drop 1 | Gap | Frequency (kHz) | Amplitude | Drop Delay | Sort Setup (e.g. 1-4 way, 5 or 15mL tubes, 96-well plate etc) | Sort Precision Mode (Single Cell_PH2, Purity, Yield, etc.) | Index Sorting | Notes:

## S01 Configuration
S01: BD FACSymphony™ S6 Cell Sorter
Location: RTP 2104D
Laser | PMT | Mirror | Filter | Parameter | Other Fluorochromes | Fluorescent Proteins | Proliferation Dyes | Cell Cycle Dyes | Viability Dyes
UV | H | 405LP | 379/28 | BUV396 | BUV380 ¦ AF350
355 nm | G | 410LP | 450/50 | DAPI |  |  |  |  | Hoechst Blue
(65 mW) | F | 450LP | 515/30 | BUV496
| E | 550LP | 586/15 | BUV563
| D | 600LP | 610/20 | BUV615
| C | 635LP | 670/30 | BUV661 |  |  |  |  | Hoechst Red
| B | 690LP | 740/35 | BUV737
| A | 770LP | 820/60 | BUV805
Violet | H | 410LP | 410/28 | BV421 | Pacific Blue ¦ BV450 ¦efluor450 ¦ AF405 ¦ VioBlue ¦ DyLight 405 | BFP ¦ Azurite | CellTrace Violet | DyeCycle Violet | L/D Fix Violet ¦ Zombie Violet ¦ Sytox Blue
402 nm | G | 450LP | 470/14 | BV480 | AF 430 ¦ AmCyan ¦ Pacific Green ¦ Qdot 525 ¦ V500 ¦ VioGreen | CFP ¦ Cerulean |  |  | L/D Fix Aqua ¦ Zombie Aqua
(200 mW) | F | 550LP | 586/15 | BV570 | Qdot 565 ¦ QDot 585 ¦ Pacific Orange |  |  |  | L/D Fix Yellow
| E | 600LP | 610/20 | BV605 | Qdot 605 ¦ DyLight 405 LS ¦ eFluor 605NC ¦ eVolve 605
| D | 635LP | 670/30 | BV650 | Qdot 655 ¦ SuperBright 645 ¦ eFluor 650NC ¦ eVolve 655
| C | 690LP | 710/50 | BV711 | Qdot 705 ¦ SuperBright 702
| B | 740LP | 750/30 | BV750
| A | 750LP | 780/60 | BV786 | Qdot 800
Blue | FSC |  |  | FSC
488 nm | G |  | 488/10 | SSC
(200 mW) | F | 505LP | 515/20 | BB515 | FITC ¦ AF488 ¦ Qdot 525 ¦ Cy2 ¦ DyLight488 ¦ VB515 | GFP ¦ YFP ¦ Emerald | CFSE | DyeCycle Green | L/D Fix Green ¦ Zombie Green ¦ Sytox Green
| E | 600LP | 610/20 | BB630 | PE-eFluor610 | mCitrine
| D | 635LP | 670/30 | BB660 | PerCP
| C | 690LP | 710/50 | BB700 | PerCP-E710 ¦ PerCP-Cy5.5 ¦ PerCP-Vio700
| B | 730LP | 750/30 | BB750 | BB755
| A | 770LP | 780/60 | BB790
Green | E | 570LP | 586/15 | PE | BYG584 ¦ BYG568 ¦ AF 514 ¦ AF 555 ¦ eFluor 570 ¦ Cy3 ¦ RFP | dsRed ¦ TdTomato | CellTrace Yellow | DyeCycle Orange | L/D Fix Red ¦ Zombie Red ¦ Sytox Orange
561 nm | D | 600LP | 610/20 | PE-CF594 | PE-Texas Red ¦ ECD¦ PE-eFluor 610 ¦ Texas Red ¦ AF 568 | mCherry
(150 mW) | C | 650LP | 670/30 | PE-CY5 | BYG670 | mPlum ¦ E2-Crimson |  | PI ¦ 7-AAD | PI ¦ 7-AAD
| B | 690LP | 710/50 | PE-CY5.5 | PE-AF700
| A | 750LP | 780/60 | PE-CY7 | PE-Vio770
Red | C | 655LP | 670/30 | APC 650 | APC ¦ AF633 ¦ AF635 ¦ AF647 ¦ AF680 ¦ Cy5.5 ¦ DyLight 633 |  | CellTrace Far Red | DyeCycle Ruby | L/D Fix Far Red
638 nm | B | 690LP | 710/50 | APC-CY5.5 | AF700 ¦ APC-R700 |  |  |  | Fixable Viability 700
(140 mW) | A | 750LP | 780/60 | APC-CY7 | APC-H7 ¦ APC-Fire750 ¦ APC-eFluor 780 ¦ AF750 ¦ Cy7 |  |  |  | L/D Fix Near-IR ¦ Zombie NIR

## S01 Spread Matrix
Total Spread Matrix | APC 650-A | APC-CY7-A | Alexa700-A | BB515-A | BB630-A | BB660-A | BB700-A | BB750-A | BB790-A | BUV496-A | BUV615-A | BUV661-A | BUV737-A | BUV805-A | BV421-A | BV480-A | BV570-A | BV605-A | BV650-A | BV711-A | BV750-A | BV786-A | DAPI-A | PE-A | PE-CF594-A | PE-CY5-A | PE-CY7-A
APC 650-A | 0 | 30.3732 | 94.5183 | 0 | 0 | 53.7953 | 37.8295 | 14.9692 | 16.4872 | 33.022 | 18.3335 | 430.31 | 49.396 | 16.4891 | 0 | 0 | 0 | 22.7811 | 74.1637 | 36.74 | 43.3613 | 41.0678 | 0 | 15.4013 | 14.9125 | 59.3161 | 17.2204
APC-CY7-A | 23.9446 | 0 | 16.486 | 0 | 0 | 11.4806 | 11.7635 | 18.8643 | 26.9606 | 16.8097 | 0 | 32.9241 | 50.4045 | 125.502 | 19.831 | 0 | 121.203 | 0 | 30.402 | 21.3286 | 25.1874 | 54.4848 | 0.045894 | 7.3727 | 0 | 13.8843 | 31.4493
Alexa700-A | 71.666 | 48.1569 | 0 | 0 | 8.2208 | 20.8593 | 29.0458 | 28.5583 | 28.1536 | 9.3489 | 0 | 85.382 | 312.849 | 37.7364 | 21.8187 | 0 | 85.7293 | 10.6617 | 19.2744 | 42.0147 | 55.2167 | 104.191 | 5.2917 | 0 | 5.8814 | 14.9149 | 34.8241
BB515-A | 8.4186 | 0.038827 | 0 | 0 | 18.823 | 18.8641 | 12.5807 | 16.0307 | 0 | 30.8204 | 14.0608 | 10.3258 | 0 | 6.155 | 0 | 12.2556 | 46.1034 | 32.4769 | 14.0729 | 0 | 0 | 19.7652 | 0 | 14.5204 | 5.2916 | 0 | 0
BB630-A | 123.426 | 11.1894 | 22.0528 | 27.4376 | 0 | 107.931 | 125.073 | 76.4193 | 61.5444 | 26.9157 | 246.27 | 178.021 | 62.665 | 40.097 | 49.7777 | 0 | 1141.85 | 447.909 | 229.129 | 86.8552 | 123.325 | 227.607 | 0 | 128.812 | 84.4965 | 82.5747 | 26.3055
BB660-A | 364.86 | 34.0868 | 85.187 | 0.16859 | 21.2998 | 0 | 144.728 | 84.5343 | 48.8835 | 0 | 16.2365 | 781.407 | 79.3145 | 39.2679 | 85.0021 | 0 | 287.94 | 55.017 | 188.047 | 111.714 | 124.057 | 163.727 | 5.5944 | 0 | 6.9119 | 76.7737 | 21.4396
BB700-A | 193.941 | 27.5303 | 125.595 | 14.2405 | 29.8426 | 74.7878 | 0 | 123.055 | 76.5009 | 9.8559 | 14.9695 | 418.758 | 270.178 | 56.2465 | 110.652 | 71.0755 | 207.56 | 56.2669 | 108.644 | 141.287 | 154.226 | 318.246 | 8.1173 | 14.5776 | 0 | 34.95 | 25.1855
BB750-A | 44.0355 | 43.7127 | 38.1899 | 18.531 | 0 | 27.8152 | 128.394 | 0 | 115.846 | 69.2028 | 17.7894 | 42.6703 | 387.022 | 134.687 | 0 | 28.6299 | 249.861 | 52.2115 | 48.9425 | 53.4277 | 146.458 | 298.593 | 5.8815 | 0.17309 | 7.3747 | 0 | 23.5299
BB790-A | 27.6584 | 33.2763 | 14.97 | 17.035 | 24.2556 | 18.1509 | 48.0665 | 103.086 | 0 | 22.8026 | 13.5247 | 26.2081 | 105.444 | 367.234 | 76.5278 | 27.5742 | 176.881 | 0 | 44.7529 | 15.5615 | 59.7792 | 304.866 | 0 | 8.4167 | 0 | 17.9235 | 21.4388
BUV496-A | 32.4726 | 0 | 11.4806 | 37.394 | 21.6763 | 0.30386 | 24.5353 | 0 | 7.3721 | 0 | 24.3795 | 106.851 | 21.7981 | 22.2744 | 0 | 12.3772 | 76.2165 | 0 | 31.6809 | 19.6768 | 0 | 0 | 14.9125 | 45.2363 | 0 | 0 | 0
BUV615-A | 120.345 | 8.0151 | 25.1531 | 9.3803 | 201.398 | 37.2081 | 33.8013 | 24.7288 | 21.215 | 58.1041 | 0 | 681.771 | 134.04 | 44.9183 | 63.5612 | 0 | 433.013 | 272.871 | 51.8365 | 28.2364 | 30.1738 | 83.7 | 17.6451 | 481.646 | 303.034 | 67.1721 | 33.273
BUV661-A | 394.022 | 22.9603 | 77.1731 | 9.6056 | 0 | 37.8029 | 26.9556 | 17.2206 | 13.5836 | 21.7861 | 12.3777 | 0 | 123.568 | 27.9253 | 0 | 0 | 129.558 | 0 | 25.9525 | 22.2721 | 30.9964 | 64.9035 | 10.6614 | 0 | 14.3499 | 37.737 | 16.2853
BUV737-A | 11.7654 | 43.9997 | 93.0647 | 9.6063 | 11.1902 | 19.5096 | 48.174 | 58.8341 | 29.6062 | 24.6346 | 13.5833 | 23.1546 | 0 | 74.8029 | 28.779 | 0 | 6.7009 | 0 | 0 | 24.6276 | 45.8188 | 95.1344 | 0 | 6.9115 | 0 | 0 | 18.5113
BUV805-A | 20.8937 | 31.833 | 9.2552 | 11.1939 | 17.412 | 0 | 12.5742 | 0 | 13.583 | 9.66 | 18.3796 | 1.9273 | 23.9468 | 0 | 0 | 0 | 92.1349 | 15.1893 | 23.2654 | 0 | 0.25388 | 25.696 | 12.5095 | 8.6097 | 0 | 0 | 14.3499
BV421-A | 0.089178 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.061164 | 0 | 5.8813 | 0 | 6.4173 | 0 | 0 | 31.4756 | 0 | 0 | 8.9803 | 6.912 | 0.095821 | 18.1976 | 0 | 0 | 7.7007 | 6.1547 | 0.032167
BV480-A | 7.8076 | 0 | 8.1171 | 13.5841 | 8.2198 | 0 | 0 | 0.061164 | 0.066614 | 39.2955 | 13.0889 | 27.6569 | 11.5517 | 5.5944 | 8.8282 | 0 | 129.951 | 28.1216 | 29.0895 | 14.9686 | 19.8852 | 26.17 | 0 | 0.041209 | 0 | 0 | 0
BV570-A | 0 | 5.5944 | 10.661 | 0 | 25.8335 | 17.4599 | 20.8572 | 0 | 0 | 0 | 21.7833 | 19.2514 | 18.3781 | 11.1153 | 29.7268 | 0 | 0 | 41.7387 | 35.9049 | 19.6774 | 23.2876 | 57.4641 | 0.032167 | 24.6245 | 13.584 | 20.6557 | 10.1874
BV605-A | 80.1048 | 0 | 29.2726 | 26.1133 | 68.2888 | 43.645 | 40.784 | 31.4441 | 15.405 | 35.3427 | 123.089 | 210.366 | 82.9324 | 56.0406 | 119.907 | 95.5594 | 725.971 | 0 | 184.415 | 74.1927 | 110.018 | 184.81 | 0 | 68.8102 | 40.4328 | 71.2773 | 28.8172
BV650-A | 204.721 | 18.38 | 67.4884 | 23.135 | 24.1974 | 27.8062 | 15.8277 | 22.2353 | 21.7829 | 47.2598 | 43.5966 | 700.05 | 120.322 | 49.7574 | 133.379 | 0 | 443.599 | 104.074 | 0 | 123.304 | 157.067 | 226.835 | 0 | 31.6509 | 24.085 | 51.022 | 18.333
BV711-A | 44.6478 | 46.7089 | 115.834 | 9.6982 | 9.5208 | 22.0595 | 61.2721 | 55.5812 | 29.1918 | 36.0195 | 0.19771 | 128.691 | 744.527 | 102.135 | 150.497 | 14.5091 | 65.9691 | 18.6936 | 73.5347 | 0 | 275.646 | 469.161 | 6.1553 | 22.7814 | 0 | 0 | 18.3326
BV750-A | 22.675 | 28.1246 | 28.8148 | 18.8319 | 0 | 14.2359 | 22.6742 | 45.1254 | 22.9609 | 0 | 0 | 0 | 363.371 | 103.851 | 89.1325 | 0 | 0 | 0 | 0 | 27.6605 | 0 | 263.161 | 0 | 0 | 9.255 | 0 | 16.6363
BV786-A | 29.8308 | 23.6669 | 11.9719 | 0 | 12.047 | 8.2157 | 18.1606 | 23.4234 | 17.7848 | 32.5236 | 13.1538 | 34.4935 | 109.463 | 222.785 | 92.2129 | 0 | 72.2931 | 36.3062 | 28.886 | 21.7827 | 46.8641 | 0 | 5.2917 | 11.763 | 0 | 0 | 16.9785
DAPI-A | 0 | 31.5161 | 26.6203 | 75.9416 | 54.9878 | 0 | 0 | 37.5114 | 0 | 314.476 | 36.0794 | 127.501 | 0 | 82.2488 | 87.5334 | 0 | 658.158 | 0 | 0 | 0 | 0 | 0 | 0 | 87.0263 | 30.1889 | 68.5531 | 0
PE-A | 51.3581 | 6.4174 | 6.6688 | 8.2198 | 114.612 | 44.9384 | 61.4545 | 27.5346 | 18.8656 | 22.4419 | 64.8987 | 73.4188 | 26.6813 | 15.4551 | 77.3811 | 0 | 517.15 | 112.386 | 94.9632 | 24.7317 | 42.361 | 76.811 | 0 | 0 | 42.9908 | 62.9295 | 19.2097
PE-CF594-A | 151.076 | 15.4026 | 28.1227 | 13.0876 | 315.95 | 167.538 | 141.134 | 100.386 | 56.2729 | 64.7962 | 184.361 | 160.027 | 49.8529 | 33.0599 | 63.1911 | 0 | 849.487 | 349.236 | 200.687 | 92.7186 | 106.429 | 181.865 | 0 | 131.266 | 0 | 185.737 | 65.0025
PE-CY5-A | 650.877 | 39.2708 | 122.712 | 0 | 34.6432 | 389.631 | 267.102 | 152.659 | 87.7652 | 19.2274 | 10.268 | 1306.61 | 78.6514 | 41.4996 | 121.113 | 49.2493 | 327.934 | 99.7909 | 338.66 | 109.855 | 147.362 | 243.95 | 9.6898 | 21.7807 | 17.6451 | 0 | 92.2207
PE-CY7-A | 12.3765 | 18.2867 | 7.7008 | 11.9718 | 10.8878 | 11.553 | 31.7269 | 62.1585 | 108.988 | 14.2402 | 6.4171 | 0.23502 | 42.0105 | 118.892 | 24.3006 | 20.8992 | 0 | 0 | 23.9437 | 10.5879 | 30.3724 | 154.618 | 5.2917 | 19.883 | 11.1155 | 6.6693 | 0